###### University of Utah

Biosafety Manual and Standard Operating Procedures

Biosafety Level 2 Enhanced Laboratory

of

Dr. (name of PI)

(title)

(Institute or department and building, room location)

(date developed)

(annual review/revision date)

This document serves as a template for a Biosafety Manual and Standard Operating Procedures for BSL-2 laboratories at the University Of Utah. . It has been developed from earlier model Manuals and Standard Operating Procedures (SOPs), the University of Utah Exposure Control Template, Safety Manuals and SOPs developed at other Universities, including Arizona State University, University of Washington, University of California Los Angeles, University of California Berkeley, and Duke University, and templates provided by the American Biological Safety Association.

This model reflects the requirements of the University of Utah Institutional Biosafety Committee for appropriate BSL-2 enhanced Biosafety manuals and SOPs.  PIs should customize this document for their laboratory requirements and incorporate into their Institutional Biosafety Committee (IBC) registration in BioRAFT.

University of Utah

Biosafety Level 2 enhanced Standard Operating Procedures

(Building, room number)

Table of Contents

- Purpose

- Background

B.1. Biosafety Level 2 Enhanced

B.2. Black Box Warnings

- Introduction to General Safety and Training for the BSL-2-enhanced Laboratory

C.1. Required Training

C.2. Administrative Procedures

C.3. General Laboratory Safety

C.4. General Biosafety Cabinet Safety

C.5. Spills and Exposure Procedures

C.6. Post Exposure Procedures

C.7. Laboratory Signage

D.	Standard Operating Procedures

D.1. Containment Requirements

D.1.a Laboratory Entry/Exit

D.1.b Specimen Transport

D.2. Personal Protective Equipment

D.3. Respiratory Protection

- Proper Use and Maintenance of Equipment

E.1. Biological Safety Cabinets

E.2. Incubators

E.3. Centrifuges

E.4. Autoclave

- Waste Disposal and Housekeeping

F.1. Waste Disposal

F.2. Laundry

F.3. Warning Labels

- Emergency Equipment

G.1. Fire Extinguishers

G.2. Telephone

- Repair and Service

- Experimental Procedures

J.	Emergency Phone Numbers and Procedures

J.1. Emergency Phone Numbers

J.2. General Emergency Procedures

- Shipping and Transport

Appendices

(These are provided as examples but incorporate information that the IBC requires and reviews prior to Registration approval)

- Spill and Exposure Procedures

- Post Exposure Procedures

- Spills and Exposure Event Reporting Procedure

- SOPs for Working with Human Cells

- Lab Specific SOPs (viral production, viral titering, replication competent virus testing, transduction of cells, animal injections, animal care, post exposure, transport of materials to CCM, etc.)

Lentivirus SOP provided as example

- Testing Lentiviral Vector Stocks for Replication Competent Virus

- Viral Vector Plasmid Maps

- Gene editing Questionnaire

- Biological Toxins SOP Template

- Agent Summary Statement

- Post-Exposure Fact Sheet

- Use and Disposal of Sharps

- Sharps injury Log

- Disposal of Biohazardous Waste

- Safe Handling of Cryogenic Liquids

- BMBL BSL2 Laboratory Criteria and Design

A	Purpose

The standard operating procedures (SOPs) described in this manual apply to all research staff, hosted visitors and guests, volunteers, building staff, and service staff who enter the laboratory. This manual will be reviewed annually by the Principal Investigator or Laboratory Supervisor for changes or corrections to ensure that it is accurate.

Research involving recombinant/synthetic nucleic acids, infectious agents, creation of transgenic or knockout animals, acute biological toxins, unfixed non-human primate materials, or human blood, cells, or unfixed tissues must submit a research registration through BioRAFT to the Department of Occupational and Environmental Health and Safety (OEHS) at least every three years.  This work may require prior review and approval by the Institutional Biosafety Committee, according to the NIH Guidelines for Research Involving Recombinant DNA Molecules.  Since the university receives funding from NIH, all research at the University of Utah must comply with the NIH Guidelines.

This plan was designed as a supplement to the University of Utah Biosafety Manual to meet the requirements of the Biosafety in Microbiological and Biomedical Laboratories, 5th ed.

A risk assessment of biological agents should take into account both the intrinsic hazards of the agent as well as considerations of individuals handling the agents, addressing the potential for emergent hazards from rDNA work, and hazards associated with the vectors used for insertions. Agent hazards are those risks that are intrinsic to the agent being handled such as the following:

- Capability to infect and cause disease in a susceptible human host.

- Severity of the disease

- Infectious dose

- Availability of preventative measures

- Availability of effective treatments

- How the agent is transmitted (i.e. route of exposure)

- Quantity, concentration, and total volume used

- Stability in the environment

- Zoonotic concerns

- Allergenicity

Exposure sources in the laboratory are hazards that could result in the infection of researchers or the public through work with biological agents. Some of the more common hazard considerations include the following:

- Aerosol generation (e.g. pipetting, mixing, blending, grinding, sonicating, vortexing, centrifuging, shaking)

- Manipulation with sharps

- Animal handling

- Contact with blood, bodily fluids, or other potentially infectious material

- Ingestion of agents via contaminated work areas

- Eye-splashes from liquid nitrogen storage

When performing a risk assessment of laboratory procedures, all potential routes of exposure should be addressed. Most laboratory-acquired infections have resulted from inhalation of aerosols, splashes or sprays, and needlesticks. It is good practice to look for potential exposures via ingestion, inoculation, inhalation, and contamination of skin and mucous membranes and attempt to identify safer alternatives and risk mitigation strategies.

# Background

Review of this Manual is mandatory for all employees and staff working in the BSL-2 laboratory of Dr. (name of PI). The Manual has been customized to provide lab-specific provisions to identify and protect all personnel who may be at risk of exposure.  This Manual must be updated at least annually, and whenever there are changes in laboratory procedures that may change a worker’s exposure.  A copy of this Manual must remain in the lab, and must be accessible to lab personnel.

This Manual applies to work with (name of agent(s)).  .

- Risk Assessment

Insert Description Here:  include consideration of parent and recombinant agent pathogenicity, virulence, infectious dose, route of transmission and exposure sources, genetic alterations that alter the pathogenicity, host range, availability of vaccines and prophylactic measures, and stability.

- Signs and Symptoms of Infection

Insert Description Here

- Prophylaxis and Treatment Options

Insert Description Here

- Methods of Disinfection and Inactivation

Insert Description Here

Individuals who are immunocompromised (e.g. individuals with immunologic conditions or are being treated with immunosuppressive agents, pregnant woman, individuals with AIDS) should be advised of the risk of working with this agent and it is recommended that they consult with a healthcare provider.

B.1	Biosafety Level 2 Enhanced

Biosafety Level 2 enhanced (BSL-2 enhanced, sometimes referred to as BSL-2+) is a term frequently used to describe laboratories where work with microorganisms is conducted in a BSL-2 laboratory with biosafety practices and procedures that are typically found at BSL-3. BSL-2 enhanced is not a recognized containment level in biosafety guidance documents such as the Centers for Disease Control and Prevention’s (CDC) Biosafety in Microbiological and Biomedical Laboratories (BMBL) or the National Institutes of Health’s (NIH) Guidelines for Recombinant DNA Research.  However, the NIH’s Biosafety Considerations for Research with Lentiviral Vectors refer to “enhanced BL2 containment”.

The use of BSL-3 practices and procedures in a BSL-2 laboratory allows for research work with microorganisms, including viral vectors, to take place in an environment where the safety practices are enhanced over and above the practices required at BSL-2. BSL-enhanced is not appropriate for RG3 pathogens that are infectious via the inhalation route. At a minimum, such pathogens must be utilized in a BSL-3 laboratory with BSL-3 practices.

There is no standardized list of microorganisms, viral vectors or research projects that should be conducted at BSL-2 enhanced. Each decision to use BSL-3 practices in a BSL-2 laboratory must be made via the risk assessment process. The risk assessment serves to guide the selection of appropriate biosafety levels and microbiological practices, safety equipment, and facility safeguards that will contribute to preventing a laboratory exposure.

As outlined in the BMBL, the steps of the risk assessment process include:

1. 	Identify agent hazards and perform an initial assessment of risk.

2. 	Identify laboratory procedure hazards.

3. 	Make a determination of the appropriate biosafety level and select additional precautions as indicated by the risk assessment.

4. 	Evaluate the proficiencies of staff regarding safe practices and the integrity of safety equipment.

5. 	Review the risk assessment with a biosafety professional, subject matter expert, and the Institutional Biosafety Committee (IBC).

The risk assessment process must be applied to every new or revised research project. Examples of when BSL-2 enhanced may be appropriate include:

• 	Viral vectors encoding oncogenes, toxins or genes of unknown function, or that produce products that reduce or knockout the expression of tumor suppressors.

• 	Drug-resistant Risk Group Two (RG2) bacteria such as methicillin resistant Staphylococcus aureus (MRSA).

• 	Low titer and small volumes of Human Immunodeficiency Virus (HIV), an RG3 agent.

• 	High concentrations (>106 PFU/mL) of RG2 viruses.

• 	Work with greater than 10 liters of an RG2 agent.

Significant differences from standard BSL-2 practices can include:

- A sharps policy is implemented and sharps (e.g., glass Pasteur pipettes, needles) are not allowed. Plasticware is substituted for glassware.

- All work is performed in a BSC.

- Lab personnel must participate if medical surveillance is required per direction of IBC and Occupational Medicine physician.

- The PI must provide training to lab personnel who may not have experience working with the materials to be used with BSL-3 practices, e.g., an apprentice program may be established for personnel where they shadow more experienced personnel and are not allowed to work independently until they demonstrate proficiency.

- Create a “Spill Kit” and store within the lab. Consider a yearly “shut down” for a few days to accommodate servicing and maintenance activities.

- Disposable, solid-front, fluid-resistant gowns rather than front opening gowns.

- Two pairs of gloves.

- Safety glasses with side-shields should be worn while in the lab.

- Other BSL3 practices, based on the risk assessment by the PI and IBC.

B.2.	Black Box Warnings

The IBC has adopted a series of “Black Box Warnings” for agents that would typically require BSL2 enhanced containment. This phrase is required by the Food and Drug Administration as the strictest warning put in the labeling of prescription drugs or drug products when there is reasonable evidence of an association of a serious hazard with the drug. It is also used by agencies to highlight a higher than standard risk.  If the laboratory is proposing any of the following agents, please add the following statement to the Biosafety Manual, using a bolded and enlarged font, ensuring it is clearly identifiable by laboratory personnel:

- Black box warning (lentivirus expressing oncogene): The viral vectors used in this research pose a greater risk to laboratory workers than typical viral vectors. This research includes one or more lentivirus/retrovirus vectors that infect human cells efficiently, and that express proven or potential oncogenic proteins or other proteins that alter the cell cycle. Accidental human exposure to these vectors may increase the risk of tumor formation.

- Black box warning (lentivirus expressing small RNA molecule): The viral vectors used in this research pose a greater risk to laboratory workers than typical viral vectors. This research includes one or more lentivirus/retrovirus vectors that infect human cells efficiently, and that express small RNA molecules intended to reduce or eliminate expression by targeted genes. These molecules may inhibit expression of tumor suppressors through a targeted or off-target effect. Accidental human exposure to these vectors may increase the risk of tumor formation.

- Black box warning (HIV pseudovirus): The viral vectors used in this research pose a greater risk to laboratory workers than typical viral vectors. This research includes a pseudovirus vector system with a low genetic barrier to the generation of replication-competent HIV virus. Accidental human exposure to these vectors could result in HIV infection.

- Black box warning (HIV infectious molecular clone): The viral vectors used in this research pose a greater risk to laboratory workers than typical viral vectors. This research includes a plasmid capable of producing replication-competent HIV virus if expressed in a human cell. Accidental human exposure to these vectors could result in HIV infection.

- Black box warning (VSV-G pseudotyped HIV infectious molecular clone): The viral vectors used in this research pose a greater risk to laboratory workers than typical viral vectors. This research includes a viral vector capable of producing HIV infection in the event of an accidental exposure. Furthermore, the viral vector has the VSV-G envelope protein on its surface. This protein makes the vector capable of infecting a wider range of human cells, so may make it more likely to result in infection in the event of an accidental exposure. Accidental human exposure to these vectors could result in HIV infection.

- Black box warning (CRISPR/Cas system on a single vector): The viral vectors used in this research pose a greater risk to laboratory workers than typical viral vectors. This research includes a CRISPR(clustered regularly interspaced short palindromic repeat)/Cas system capable of editing human chromosomal DNA. Both components of the system (gRNA and Cas9) are expressed on the same viral vector. This increases the risk for damage to chromosomal DNA with potential oncogenic effects. Accidental human exposure to these vectors may increase the risk of tumor formation.

C. 	General Safety and Training for the BSL-2-enhanced Laboratory

C.1.	Training

All personnel who work in the laboratory must receive adequate instruction from their supervisor prior to beginning work.  Some training is required annually.  Each lab will require different trainings (edit as applicable). The minimum requirements for qualification to work in the (name of PI) BSL-2-enhanced lab are:

- Annual OEHS Bloodborne Pathogen training,

- Annual OEHS Biosafety Level 2 Training

- Annual OEHS Animal Biosafety Level 2 Training

- Annual Laboratory Specific Training (Risks associated with the hazards/agents used in the lab, SOPs, Spill and Exposure Procedures)

- Hands on training with BSL-3 practices. An apprentice program may be established for personnel where they shadow more experienced personnel and are not allowed to work independently until they demonstrate proficiency.

- Shipping Training: Training is required to commercially transport infectious materials, as well as shipping anything on dry ice.  Contact EHS to schedule a training class

- Annual Review of this Manual.

- (List as many as you need)

This document will provide the basis of training in conjunction with OEHS training and hands-on training provided by the PI or Lab Manager.

Dr. (name of PI) will provide information and arrange for training at the time of an individual's initial assignment to the lab. Dr. (name of PI) will arrange for refresher training at least annually and when there are any changes in processes or procedures. Documentation of training will be uploaded to the laboratory registration in BioRAFT.

For more information concerning training, also see the University of Utah Biosafety Manual on the OEHS web site (http://d2vxd53ymoe6ju.cloudfront.net/wp-content/uploads/sites/4/20160817141055/uu-ehs-biosafety-manual-revised-2015-08.pdf).

C.2. 	Administrative Procedures

It is the responsibility of each lab member to carefully consider every action taken in the BSL-2-enhanced lab and its potential impact on possible exposure or contamination, and to follow established Standard Operating Procedures (SOPs) and protocols diligently and without variance.

All employees will read and adhere to the University of Utah Biosafety Manual and to this SOP Manual for Dr. (name of PI) laboratory. All employees will use pertinent sections in the Biosafety in Microbiological and Biomedical Laboratories, CDC-NIH, 5th edition as a guideline and reference (http://www.cdc.gov/biosafety/publications/bmbl5/index.htm).

All employees will attend the University of Utah OEHS courses in Bloodborne Pathogen/Biosafety Level 2 Training annually. Records of certification will be kept on file by OEHS.

All employees working in the lab will be will be offered (vaccination available?).

No employee will be trained to work in the lab without the express permission of Dr. (name of PI).

New SOPs and protocols must be approved by the PI before initiation. Changes to SOPs and protocols that include new agents, risks or change the basis of IBC approved protocols must be submitted for review by the IBC.

Current SOPs and protocols will be reviewed and/or revised by Dr. (name of PI) or Laboratory Manager every ( ) months.

C.3. 	General BSL-2-enhanced Laboratory Safety

# Work with the agent will be performed in (insert building and room number).

- Laboratory employees must immediately notify the laboratory manager or PI in case of an accident, injury, illness, or overt exposure associated with laboratory activities.

- No eating, drinking, smoking, handling contact lenses, or applying cosmetics in the lab at any time.

- No use of cell phones/iPODs etc., while in the BSL-2-enhanced lab.

- No animals or minors (persons under the age of 18) will be allowed to enter the lab at any time.

- Food, medications, or cosmetics should not be brought into the lab for storage or later use. Food is stored outside in areas designated specifically for that purpose.

- No open-toed shoes or sandals are allowed in the laboratory.

- No bare legs or arms.

- Personnel working with virus must wear a lab coat (or gown) with tight cuffs, safety glasses or facemask, and two pairs of gloves. For work with high risk agents transmissible by inhalation, personnel should wear an N95 respirator (see Section D.3.). Lab coats must be left in the lab.  Upon entering the tissue culture facility, personnel must immediately don a gown or laboratory coat with long sleeves. Optional disposable sleeves with tight cuffs are available for personnel to use when using a lab coat with loose cuffs.

- All skin defects such as cuts, abrasions, ulcers, areas of dermatitis, etc. should be covered with an occlusive bandage.

- Mouth pipetting is prohibited; mechanical pipetting devices are to be used at all times.

- All procedures are to be performed carefully to minimize the creation of splashes or aerosols. Aerosols containing RG2/BSL-2 agents pose an increased risk of exposure to personnel. Solutions containing RG2/BSL-2 agents will be enclosed in capped containers during agitation (vortexing, shaking, centrifugation, etc.). If agitation is performed outside of a safety cabinet, two levels of containment are required (i.e. capped tube enclosed in plastic bag or Tupperware). After agitation, containers will only be opened in a biosafety cabinet to prevent exposure to personnel from aerosols. Human cells/tissue will be homogenized inside a biosafety cabinet. Centrifugation of hazardous materials must be performed using centrifuge safety caps or sealed rotors: these should loaded and unloaded in the BSC.

- Centrifuge materials containing infectious agents in durable, shatter resistant, closable tubes. Use a centrifuge with sealed rotors/heads or screw capped safety cups. After centrifugation, open the tubes within a BSC.

- Cover countertops where hazardous materials are used with plastic backed disposable paper to absorb spills and dispose of them daily or following a spill.

- Wipe work surfaces with an appropriate disinfectant (name of disinfectant) after experiments and immediately after spills. Follow manufacturer instructions for contact time.

- Decontaminate all contaminated or potentially contaminated materials by appropriate methods before disposal

- Follow all manufacturer’s instructions and SOPs when using any of the laboratory equipment.

- Hand washing with soap is required (i) after handling materials involving organisms containing recombinant DNA molecules including viruses, (ii) after handling animals, (iii) after working with human cells, and (iv) whenever exiting the laboratory.

- Sharps and glass (e.g., Pasteur pipettes) are generally prohibited in BSL-2-enhanced procedures, unless absolutely unavoidable. Consultation must be performed with OEHS to assess risks and changes to the BSL-2-enhanced protocol prior to their use. Clear justification must be provided to the IBC justifying the use of sharps prior to IBC approval. Although plastic pipette tips are considered puncture hazards they may be used in BSL-2-enhanced protocols.

- All sharps must be placed in a rigid, puncture resistant, closable and leakproof container, which is labeled with the word "Sharps" and the biohazard symbol. Food containers (e.g., empty coffee cans) are not permissible as sharps containers. All sharps must be handled with extreme caution. The clipping, breaking, and recapping of needles is not permitted. Sharps containers should not be filled more than 2/3. See Appendix 12.

- Fluorescence activated cell sorting (FACS) can produce aerosols and special precautions must be performed. BSL-2-enhanced samples that only require analysis must be fixed in 4% paraformaldehyde prior to running on a cell sorter. For sorting BSL-2-enhanced cells, permission must first be obtained from the FACS operator prior to scheduling the run. The FACS operator must be informed about the type of cells being run and the hazards of the cells.

- All cultures, stocks, and other regulated wastes are decontaminated by autoclaving or disinfection before disposal. NOTE: No biological agent-containing material should be allowed into any drain connected to the sanitary sewer system (e.g., from a sink) Materials to be decontaminated outside of the immediate laboratory are placed in a durable, leakproof container and closed for transport from the laboratory.

- Decontaminated liquid waste from humans or animals can be disposed of directly by flushing down a sanitary sewer.

- Contaminated solid waste includes cloth, plastic and paper items that have been exposed to agents infectious or hazardous to humans, animals, shall be placed in biohazard bags and decontaminated by autoclaving.  Alternatively, OEHS can dispose of biohazard waste for the lab.  OEHS can be contacted using the on-line pick up request. Double or triple bagging may be required to avoid rupture or puncture of the bags. Contaminated Pasteur pipettes and plastic pipette tips are considered sharps and need to be disposed of in a sharps container.

# C.4.	General Biosafety Cabinet Safety (edit or delete this section as appropriate)

            - Reserve time in advance to have priority for use of the BSCs. (If you use sign-up sheets, describe the process here)

            - Turn on the blower in the cabinet at least 10 minutes before placing infectious materials into the hood.

            - Check the certification sticker and all biosafety cabinet monitors to verify that the biosafety cabinet is working properly. Biosafety cabinets must be certified prior to use. A qualified outside contractor must certify these cabinets annually. Check the certification sticker on the front of the unit to verify your biosafety cabinet’s condition. If the re-certification date has passed contact OEHS.

            - The biosafety cabinet air flow monitor should be checked to assure proper operation of the cabinet before placing any materials into it. Readings indicate relative pressure drop across the HEPA filter. Higher readings may, therefore, indicate filter clogging. Zero readings may indicate loss of filter integrity. In either of these cases, notify the Laboratory Manager or PI and OEHS.  University of Utah Facilities Management do not perform maintenance on biological safety cabinets.  If the BSC needs to be serviced, contact OEHS.

            - Gloves must be worn at all times.

            - Prior to beginning work, the BSC should be decontaminated. Don appropriate PPE (rear closing, fluid resistant lab coat, gloves, eye protection). Clean the inside surfaces of the BSC with (name of disinfectant) and follow with water (if using bleach). DO NOT put head inside the cabinet. To reach the back of the cabinet use an extension, such as a Swiffer handle.

            - Let blower run for 10 min to filter the cabinet air of any particulates.

            - DO NOT disrupt the airflow through the hood by placing ANY item on the grills or by opening the door to the corridor. Disrupting the airflow into the front grill allows contaminated air from inside the cabinet to blow into the lab or directly at the person sitting at the cabinet. It also allows non-sterile air from the room to blow into the biosafety cabinet over the experiments.

            - Organize the work surface for a clean-to-dirty work flow. Place clean pipets, flasks, and sterile media bottles at one side of the cabinet; place discard or kill pans containing disinfectant, biohazard waste containers, used flasks, spent cultures, and other wastes on the other side.

            - While working, keep all material and perform work at least 4 inches back from the front opening of the cabinet, and minimize rapid movements or activity.

            - In general, the interior of the hood should be considered to be a contaminated zone, even though every effort is made to keep the surfaces clean, as is consistent with accepted good microbiological practice and sterile technique.

            - After manipulating infectious agents, make sure all containers are tightly closed.

            - Plastic pipettes with a cotton plug shall be used for pipetting liquids containing viral particles. The electric pipettor shall be fitted with a 0.2 µm filter to prevent aerosol-based contamination.

            - A beaker or discard pan, containing a freshly prepared 1:10 solution of commercial bleach, shall be placed inside the biosafety cabinet during the cell culture work.

            - After pipetting liquid containing viral particles, the dilute bleach solution in the beaker shall be pipetted up and down the full length of the pipette or left in the pan. Serological pipettes and tips should be placed in a in a puncture resistant sharps container or other approved receptacle.

            - After decontamination, pipette tips shall be removed from the pipettor and temporarily left in the beaker containing bleach in the biosafety cabinet.

            - At the completion of the work, all materials to be removed from the biosafety cabinet must be decontaminated.  Describe methods for decontamination.

            - At the completion of the work, the beaker containing the plastic tip pipettes shall be removed from the biosafety cabinet.  Pipettes tips shall be lifted out of the beaker, the bleach solution allowed to drain back into the beaker, and the pipette tips placed in a puncture resistant sharps container or other approved receptacle.  NOTE plastic pipette tips and serological pipettes are treated as sharps.

            - Small volumes of liquid waste containing viral particles shall be collected in a beaker containing undiluted bleach inside the biosafety cabinet. The final concentration of bleach should be at least 10% of the final volume. After completing work, wait at least 30 minutes before disposing down the drain.

            - Large volumes should be collected by vacuum aspiration into a flask containing an appropriate disinfectant, such as Wescodyne or bleach, up to 1 or 10% of the volume of the flask, respectively. NOTE: No biological agent-containing material should be allowed into any drain connected to the sanitary sewer system (e.g., from a sink).

              - Bleach in the vacuum traps must be changed at least twice per week or when the flask is half full. Wait at least 20 minutes after finishing work to discard waste.

              - Wescodyne is more stable and can be used for up to 3 months before it is discarded. Wescodyne-treated waste must be transferred to a sealed container marked “Unwanted materials”. When full or 6 months after your start date (whichever happens first), arrange pickup by OEHS through the LMS website.  NO DRAIN DISPOSAL.

              - The flask should be placed in a secondary container to prevent it from tipping over, be labeled with a biohazard sticker and the vacuum line should be protected by a hydrophobic (HEPA) filter. The vacuum filters must be replaced if clogged or if liquid makes contact with the filter. Examples include Whatman Vacu-guard and Pall Gelman Vacushield in-line disk filters.   Used filters should be placed in the biohazard waste.

            - Turn off the house vacuum when not in use.

            - Clean the inside surfaces of the BSC with (name of disinfectant) after completion of work, and follow with water (if using bleach).

            - Allow the blower to run for at least 10 minutes following use.

            - The UV light is turned on between procedures (at least 30 minutes). UV lights must be turned off before work begins in the hood. Do not look directly at UV lights as this can cause eye damage.

During decontamination, a sign shall be placed stating that the biosafety cabinet is being decontaminated and shall not be used. Also, it shall be stated that the user (with contact information) shall be contacted if there are any questions or concerns.

UV light is effective only for decontaminating clean, solid surfaces with which it comes in contact. It is not effective in decontaminating the cabinet air flow. UV light is not effective against bacterial spores. UV germicidal light tubes should be replaced frequently (at least every 6 months for biosafety cabinets in use on a daily basis) to assure that they are emitting light at 254 nm and at an intensity appropriate for decontamination. Due to concerns over the effectiveness of these lights and the risks to individuals in the room, some Institutions, such as the NIH, have banned their use in BSCs.

NOTE: Any use of volatile solvents, such as absolute ethanol, should be kept to a minimum or done elsewhere. Dangerously high levels of volatile vapors can accumulate inside the cabinet and pose a threat of fire or explosion.

C.5.	Spills and Exposure Procedures

See Appendix 1

C.6.	Post-exposure Procedures

See Appendix 2

C.7.	Laboratory Signage

Permanent yellow warning signs containing the biohazard symbol and BSL2 designation are posted at the entrance to rooms where Risk Group 2 agents are used or stored.  This system is intended for rooms where the use of BSL2 containment is expected to continue over a long time period without frequent total room decontamination. This door sign outside of the laboratory is posted and managed by OEHS.  Biosafety level 2 labs should have the following permanently-affixed decals: Biohazard symbol, “BSL-2” designation, and Exit Requirements, including reminders to remove PPE and wash hands prior to exiting.  The types agents worked with in the lab are listed but not specifically identified.  If any changes need to be made to the sign, including emergency contacts, please log in and update the information in BioRAFT. EHS will post any changes to the sign after IBC approval.

# D.	Standard Operating Procedures	 (edit or delete this section as appropriate)

D.1. 	Containment Requirements

D.1.a - 	Laboratory Entry/Exit

Entering the lab to begin work

A. Put on protective PPE (see Section D.2.).

B. Gather all materials for the experiment.

Exiting Laboratory

Before exiting the lab, be sure that the hood and work area are clean, all contaminated waste materials are disposed of properly, and stocks have been returned to the proper storage area. Remove PPE and wash your hands.

D.1.b 	Specimen Transport

Transport of biological materials to another building or lab within the same building should be done in a sealed, leakproof container labeled with the universal biohazard sticker. If the samples are infectious, use a secondary container and label it with the contents and a contact person/phone number.

D.2.	Personal Protective Equipment (PPE)

The following personal protective equipment (PPE) will be required when working in the BSL-2-enhanced laboratory:

- Rear closing, fluid resistant disposable or onsite-laundered lab coat: if sleeve cuffs are not tight, disposable sleeve covers are required

- Gloves (2 pairs)

- Safety glasses or face shield

- (describe additional PPE requirements)

Staff should not have exposed arms, legs or feet. Shorts and open toed shoes or sandals are not permitted in the BSL-2-enhanced laboratory.

# D.3.	Respiratory protection (edit or delete this section as appropriate)

In a setting where there is a high risk of exposure to an agent that is contracted through aerosolization, employees will be required to wear an N95 respirator. Medical surveillance, a written respiratory protection plan, and fit testing/training must be conducted before personnel wear N95 respiratory protection.  Contact OEHS to schedule fit testing.

# E.	Proper Use of Equipment (edit or delete this section as appropriate)

# E.1.	Biological Safety Cabinets

See Section C.4

# E.2.	Incubators

E.2.a	Incubators (these must be written specific to your lab)

A. Incubators are normally set at 37o C.

B. Temperature should be checked each day by all users.

C. Operation manuals are located (describe where these are).

D. If an alarm is sounding, check the panel for the identifying blinking light.

1. If there is no obvious reason for the alarm, contact the Lab Manager or PI.

2. The “CO2 Low” (or High) message indicates a deviation from 5% CO2. Check the hose from the wall to the unit.

3. The “tank farm” must be checked for empty tanks once/week.

E. Decontaminate incubators at least every (insert length of time and method).

E.2.b	Water bath

The water bath should be monitored for water level, and filled with distilled water only. To prevent growth of any organisms, water should be treated with (name of disinfectant).

E.3.	Centrifuges

E.3.a.	Potentially hazardous material should only be centrifuged in buckets with sealed lids

E.3.b	Centrifuge buckets should be loaded and unloaded in the BSC.

E.3.c	If there is a spill or breakage during centrifugation, stop the centrifuge, and all personnel should evacuate the room for at least 30 minutes.  Upon re-entry personnel should don PPE, including eye protection and/or face shield for high risk pathogens spread by aerosolization personnel should wear an approved respirator (see E.5.). After opening the centrifuge buckets should be transferred to the BSC. The centrifuge and buckets should be disinfected by cleaning with (disinfectant) and then wiped down with 70% ethanol.

E.3.d.	(describe lab specific procedures): See Appendix 1.

E.4.	Autoclave

(describe procedures)

(Add similar lists for each type of equipment used in the BSL-2-enhanced facility, such as microplate reader, etc.)

- Waste Disposal and Housekeeping

F.1.	Waste Disposal (See Appendices 12 and 14 for detailed instructions)

F.1.a	Contaminated waste is placed in containers which are closable, constructed to contain all contents and prevent leakage, appropriately labeled or color-coded (see Labels), and closed prior to removal to prevent spillage or protrusion of contents during handling.

F.1.b.	Sharps containers will be disposed of by filling out the on-line pick up request at www.ehs.utah.edu.  OEHS personnel will pick up full containers and replace them with empty containers.

F.1.c	Other regulated waste must be autoclaved before being disposed of in the municipal waste stream.  Alternatively, OEHS can dispose of biohazard waste for the lab.  OEHS can be contacted using the on-line pick up request.

F.1.d.	Contaminated sharps are discarded immediately or as soon as possible in containers that are closable, puncture-resistant, leakproof on sides and bottoms, and labeled or color-coded appropriately. Sharps disposal containers are available through the OEHS on-line pick up request at www.ehs.utah.edu.

F.1.e.	Bins and pails (e.g., wash or emesis basins) are cleaned and decontaminated as soon as feasible after visible contamination.

Broken glassware that may be contaminated is picked up using a mechanical device such as tongs, forceps, or a brush and dustpan.

F.2	Laundry

The following articles, if contaminated, will require laundering:

The University of Utah School of Medicine can be used to clean contaminated clothing and other articles that require laundering.  Linen Services can be found in the AA120 level of the School of Medicine, 801-581-2200.

The following laundering requirements must be met:

- Handle contaminated laundry as little as possible, with minimal agitation.

- Place wet contaminated laundry in leak-proof, labeled or color-coded containers before transport to the University Hospital Laundry.

F.3.	Warning Labels

F.3.a	The universal biohazard symbol will be used to mark regulated waste containers, refrigerators containing biological agents, blood or Other Potential Infectious Material (OPIM), contaminated equipment, specimen containers, and specimen transport containers.  The universal biohazard symbol must be located at the entrance to all laboratory rooms where blood or OPIM is used.  Red bags can be used as a substitute for bags marked with the universal biohazard symbol for regulated waste only.

F.3.b.	(Name of Individual) will ensure warning labels are affixed or red bags are used as required if regulated waste or contaminated equipment is brought into the facility.

# Emergency Equipment (edit or delete this section as appropriate)

G.1	Fire Extinguisher, located (location of fire extinguisher).

G.1.a. 	Operation

- Fire extinguishers should be used only if the fire is small and confined to one small area! USE JUDGEMENT IN THIS! DO NOT CREATE A LIFE-THREATENING SITUATION WHILE TRYING TO EXTINGUISH A FIRE!

- To operate, pull the pin to release the handle.

- Stand at a safe distance from the fire (as directed on the fire extinguisher).

- Aim the nozzle at the base of the fire, squeeze the handle to discharge the agent, and sweep completely left and right until a few seconds after seeing no fire.

G.1.b. 	Maintenance

Fire extinguishers are inspected annually by OEHS. Check the gauge periodically to ensure operational status. Call OEHS at 581-6590 if you have any questions.

G.1.c. 	Training

OEHS offers monthly fire extinguisher trainings to cover elements necessary for a fire and proper use of a fire extinguisher.  Laboratory employees working with flammable and/or combustible materials are encouraged to enroll in this training.  The following link will direct you to the OEHS website to register for the training, http://training.ehs.utah.edu/class_details.jsp?offeringId=88.

G.2.	Telephones are located (describe locations).

# Repair and Service (edit or delete this section as appropriate)

(describe lab policy for decontaminating equipment before having service done on it)

# Experimental Procedures

(include major protocols for lab procedures or refer to material in appendix if described there)

See Appendix 5 for examples

- Emergency Phone Numbers and Procedures

# J.1 Emergency Phone Numbers

| Fire and Medical Emergencies | 911 |
| --- | --- |
| Police | 911 |
| Laboratory PI Home Phone Office Phone Cell Phone |  |
| Laboratory Manager Home Phone Office Phone Cell Phone |  |
| Occupational Medicine | 801-213-9777 |
| Campus Police Department | 801-585-2677 |
| Department of Occupational and Environmental Health and Safety (Main Number) | 801-581-6590 |
| Biosafety Officer (OEHS) | 801-585-9325 |
| Associate Biosafety Officer (OEHS) | 801-585-3345 |

# J.2 Emergency Procedures

In responding to an emergency situation, all persons must be aware of these basic concerns:

Emergency Equipment:

telephoneS are located (insert location).

fire extinguisher is located (insert location).

eyewash is located (insert location).

SAFETY showeR is located (insert location)

Order of Priority in an Emergency:

I. SAFETY FOR THE EMPLOYEE/STUDENT(S) IN THE LABORATORY.

Notify anyone else present in the lab.

Safety for persons in the lab must be the first consideration. The circumstances of the emergency will determine the feasibility of carrying out site control.

II. DANGER OF OUTSIDE CONTAMINATION (Site Control):

Do not exit building wearing contaminated clothing or protective outerwear (gown, gloves).

Leave biosafety cabinet on.

Place covers on open containers of viable agents.

K.	Shipping and Transport

Packaging of (name of agent) for Shipping to Other Institutions

(list authorized person for shipping hazardous materials (including Dry ice)– specialized training is required for shipping hazardous materials including infectious agents, contact OEHS for more information)

NOTE: Category A Infectious Substances have to be shipped by OEHS.

Version #:

Date Created:

Created by:

Approved by:  PI Signature and Date

SOP ANNUAL REVIEW AND APPROVAL

New SOPs and protocols must be approved by the PI before initiation. Current SOPs and protocols will be reviewed by (name of PI) or Laboratory Manager every ( ) months.

| ANNUAL REVIEW LOG |
| --- |
| MONTH/YEAR | PRINTED NAME, SIGNATURE, DATE |

TRAINING

All lab personnel must review these SOPs prior to starting work in the lab and at least annually thereafter.  Copies of this record will be attached in BioRAFT laboratory registration annually.

| TRAINING LOG |
| --- |
| DATE | PRINTED NAME, SIGNATURE, DATE |

Appendices

Appendix 1: Spills and Exposure Procedures

All spills or breaks involving Recombinant DNA or Synthetic Nucleic Acid Molecules and hazardous biological materials should be cleaned up using appropriate biosafety procedures, described below. If there is any doubt about what to do, call the PI (Telephone #), or the Biosafety Officer 1-6590, or the University’s internal emergency number: 5-2677.

- Spills inside of a Biosafety Cabinet

  - Stop work.

  - If you are splashed by the material, change PPE. Always change gloves.

  - Keep the biosafety cabinet running.

  - Contain the spill by covering with paper towels (to avoid splashes or aerosols).

  - Saturate spill with XXXXXXX (fill in the appropriate decontaminant, e.g., freshly prepared 1:10 dilution of bleach). Let sit for 20 minute exposure time.

    - In the event of a spill into the drip pan/catch basin, add an equal volume of disinfectant and wait for 20 minutes to clean up the disinfected material.

  - Wipe up spill, disposing of towels in biohazard bag.

  - Wipe spill area with XXXXXXX (fill in the appropriate decontaminant).

  - Disinfect all materials used in the biosafety cabinet by wiping the surface with an XXXXXXX (fill in the appropriate decontaminant). Do not attempt to disinfect contaminated cardboard or other paper items that absorb liquid: contaminated items should be disposed of.

  - If bleach or other corrosive disinfectant used, wipe spill area and disinfected equipment with water.

  - Place all towels or absorbent materials into a designated container for biohazardous waste.

  - Remove PPE, discard disposable PPE as biohazardous waste and wash hands.

  - Run the biosafety cabinet for 10 minutes to purge the air before re-starting work.

- Spills outside of a Biosafety Cabinet

  - Stop work.

  - If you are splashed by the material, dispose of PPE and wash hands.

  - Ensure that any other people in the vicinity are notified that a spill has occurred and that the room should be evacuated. Post a “Do Not Enter” notice on the door. Notify the PI or lab supervisor.

  - If you need assistance with the spill clean-up, call OEHS (1-6590).

  - Wait 30 minutes before re-entering the room to allow aerosols to settle.

  - Assemble Spill cleanup materials and don PPE, including lab coat, eye protection and/or face shield, 2 pair of gloves, shoe covers.  If the lab coat does not have cuffed sleeves, disposable sleeve covers should be worn.

  - Contain the spill by covering with paper towels (to avoid splashes or aerosols)

  - Saturate spill with XXXXXXX (fill in the appropriate decontaminant, e.g., freshly prepared 1:10 dilution of bleach) pouring from the outside in. Avoid splashing or generating aerosols. Do not use alcohol for large spills.

    - Wipe areas around the spill that may have splatter and any reusable equipment with XXXXXXX (fill in the appropriate decontaminant).

  - Let sit for 20 minute exposure time.

  - Wipe up spill, disposing of towels in biohazard bag: if sharps may be present use tongs or a brush and pan and dispose in biohazard sharps container.

    - Work concentrically to clean up the absorbent material. Always work from the outer edge of the spill toward the center.

  - Wipe spill area with XXXXXXX (fill in the appropriate decontaminant).

  - Place all towels or absorbent materials into a designated container for biohazardous waste.

  - Remove PPE, discard disposable PPE as biohazardous waste and wash hands.

  - Remove the “Do Not Enter” sign and inform others that it is safe to re-enter the room.

  - Once the spill has been contained, complete the “SPILLS OR EXPOSURE EVENT REPORTING PROCEDURE” form (below) and have the PI send to OEHS.

- Spills Inside of a Centrifuge Contained Within a Closed Cup, Bucket, or Rotor

  - Put on lab coat, gloves, and proper eye protection prior to opening centrifuge.  Open carefully to assess the damage.

  - If the spill is contained within a closed cup, bucket, or rotor, spray the exterior with XXXXXXX (fill in the appropriate decontaminant, e.g., 1:10 dilution of bleach and/or 70% ethanol) and allow at least 20 minutes of contact time.  Remove the carrier to the nearest biosafety cabinet (BSC).  If a BSC is not available, close centrifuge; post a sign to indicate it cannot be used.  Notify the PI and Biosafety Office for assistance.

  - Gather supplies needed, such as a sharps container for broken glass and bins filled with disinfectant and place into the BSC.

  - Open the centrifuge rotor or bucket inside of the BSC. Use a mechanical device (forceps, tongs, etc.) to remove broken glass and place directly into sharps container.  Carefully remove any unbroken rubes and place into a bin filled with XXXXXXX (fill in the appropriate decontaminant) for at least 20 minutes.  Wipe carrier/bucket with disinfectant.

  - After disinfection, carrier, bucket, or rotor should be washed with a mild soap and water.

  - Spray the interior of the centrifuge chamber with XXXXXXX (fill in the appropriate decontaminant), let sit for at least 20 minutes and then wipe down.

  - Dispose of all clean-up materials (except sharps) in an appropriate biohazardous waste container.  Dispose of sharps in a biohazard sharps container.

  - Remove PPE, discard disposable PPE as biohazardous waste and wash hands.

If you are concerned that the spill is not contained within the rotor or bucket:

    - Ensure that any other people in the vicinity are notified that a spill has occurred and the room should be evacuated. Post a “Do Not Enter” notice on the door. Notify the PI or lab supervisor.

    - If you need assistance with the spill clean-up, call OEHS (801-581-6590)

    - Wait 30 minutes before re-entering the room to allow aerosols to settle.

    - Proceed with clean up as described above.

- Exposure to skin or clothing

  - Stop work.

  - Take off contaminated clothing and wash affected area thoroughly with soap and water, but not so hard the skin is abraded.

  - If necessary, exit lab area and immediately take a shower. Wash thoroughly with soap and water, but not so hard the skin is abraded.

  - Notify the lab supervisor or PI.

  - If exposed to BSL-2/RG2 (or above) agent, notify the Biosafety Officer.

- Penetrating wound

  - Stop Work.

  - Wash immediately with soap and water.

  - Notify lab supervisor or PI, who must notify the Biosafety Officer.

  - Proceed directly to Redwood Occupational Medicine Clinic, RedMed Clinic or the University of Utah Hospital emergency Room (if after 8pm).

- Eyes, or mucous membrane exposure

  - Stop work.

  - Immediately flush eyes or mucous membrane with water for 10-15 minutes.

  - Notify lab supervisor or PI, who must notify the Biosafety Officer.

  - Proceed directly to Redwood Occupational Medicine Clinic, RedMed Clinic or the University of Utah Hospital emergency Room (if after 8pm).

- Emergency Spills: Environmental Risk

  - Stop work.

  - Ensure that any other people in the vicinity are notified that a spill has occurred and that the room should be evacuated. Post a “Do Not Enter” notice on the door. Notify the PI or lab supervisor.

  - Call OEHS (801-581-6590)

  - Take appropriate precautions to limit exposure or spread of spill to other areas.

NOTE: Spill Procedures should be clearly posted in the BSL-2-enhanced suite

Appendix 2: Post-exposure Standard Operating Procedure

Complete for each of the agents/viral vectors in the laboratory

| Date: |  |
| --- | --- |
| PI Name: |  |
| Name of Agent: |  |
| Mode(s) of Transmission of Parental Agent (e.g. for recombinant lentivirus vectors describe modes of transmission of HIV): |  |
| Signs and Symptoms of Exposure/Infection with this Agent (e.g. for recombinant lentivirus vectors describe signs and symptoms of exposure/infection with HIV): |  |

| Assessment of Additional Risks |
| --- |
|  | Yes | No |
| Is this agent replication-defective? |  |  |
| Could the agent integrate into the genome? |  |  |
| Does the agent encode a transgene that could lead to an increased risk of oncogenesis? e.g. Does the transgene encode an oncogene? Could the transgene inactivate a tumor suppressor? |  |  |
| Has the agent been modified to alter drug resistance? If yes, explain: |  |  |
| Is the agent propagated in cell lines? |  |  |
| If yes, are the cell lines human or primate? |  |  |
| If yes, are there any hazards associated with the cells? |  |  |

| Vaccination and Treatment Options |
| --- |
|  | Yes | No |
| Is there a Vaccine available to protect against this agent? |  |  |
| If yes, have personnel been offered vaccination? |  |  |
| Are there Post Exposure Prophylaxis options for this agent? If yes, describe: |  |  |

Copy and complete the three tables for each type of agent/viral vector in the laboratory

Actions in the Event of an Exposure

Definition of Exposure

Direct skin, eye or mucosal membrane exposure to the agent or materials potentially containing the agent, such as tissue culture media or cells, bodily fluids from infected animals. Parenteral inoculation by a syringe needle or other contaminated sharp (needlestick), Ingestion of liquid suspension of an infected material or by contaminated hand to mouth exposure, or Inhalation of infectious aerosols.

Information for Lab Personnel

Remove exposed PPE taking care to avoid contact of unexposed areas to infectious agents on the PPE. Inform others in area about any biohazardous materials out of containment to prevent further exposure. If possible, contain with absorbent pads, decontaminate with bleach, and/or seal off the site. ALL exposed individuals should leave the area. Immediately wash affected areas with soap and water, or if exposure to eyes or mucous membranes occurred, immediately flush affected area with water for 10-15 minutes. See exposure procedures for further information. After washing, Notify lab supervisor or Principal Investigator of the exposure (PI’s 24-hour Emergency Contact Number: (XXX) XXX-XXXX). Go immediately to the RedMed Employee Health Clinic at the University Union Building or the Occupational Medical Clinic at the Redwood Health Center for medical evaluation and follow-up; contact information is below. After 5pm you will be seen by an Urgent Care Physician. After 8pm, or if the injury is serious/life threatening, go to the University of Utah Hospital Emergency Department or call an ambulance (911). Ensure that the physician is aware of all materials that were being used at the time of exposure (e.g., virus, bacteria, human tissue, animal tissue, other potentially infected material). Take a completed copy of the risk assessment and treatment options of this SOP with you! Follow up with the physician at Occupational Medicine, as requested. RedMed Employee Health Clinic 200 Central Campus Dr. Salt Lake City, UT 84112 Phone: (801) 213-3303 Hours: M-TH: 8:00AM – 5:00PM, Friday: 9:00AM – 3:30PM Redwood Health Center Occupational Medicine Clinic 1525 West 2100 South Salt Lake City, UT 84119 Phone: (801) 213-9777 Hours: M-F 8:00AM – 5:00PM After Hours Redwood Urgent Care 1525 West 2100 South Salt Lake City, UT 84119 M-F 5:00PM – 8:00PM Sat.-Sun.: 9:00AM – 8:00PM (801) 213-9700 After 8 PM Emergency Department at University Hospital (main floor northeast side of the hospital) 50 N. Medical Drive Salt Lake City, UT 84132 (801) 581-2292 Post exposure prophylaxis must be initiated as soon as possible after exposure. Inform the Healthcare Provider of any medical conditions, such as pregnancy or immunosuppression, or drug treatment that you currently have or take. The Healthcare Provider must have this information to evaluate and develop a proper post treatment evaluation. Upon returning to work, fill out the Employers First Report of Injury E1 Form. This form can be downloaded from the human resources website under “Forms” (https://www.hr.utah.edu/forms/index.php). After medical care, ensure that the incident is immediately reported to the Biosafety Officer (801-581-6590). Have the PI/Supervisor complete a “SPILL OR EXPOSURE EVENT REPORT,” using the template below, and submit to the Biosafety Officer (801-581-6590).

# Adapted from the UCLA Post Exposure SOP template (https://ucla.app.box.com/v/ehs-bio-pep-sop-general)

Maps of Occupational Medicine Clinics

UNIVERSITY OF UTAH REDWOOD HEALTH CENTER OCCUPATIONAL MEDICINE 1525 W. 2100 S. Salt Lake City UT, 84119 REDMED EMPLOYEE HEALTH CLINIC 200 Central Campus Dr. Salt Lake City, UT 84112

Appendix 3: SPILLS OR EXPOSURE EVENT REPORTING PROCEDURE

Any significant problems, violations of the NIH Guidelines, or any significant research-related accidents and illnesses must be reported to the IBC so that a report can be sent to the NIH Office of Science Policy (OSP) within 30 days. Certain types of accidents must be reported on a more expedited basis. Spills or accidents in BL2 laboratories resulting in an overt exposure must be immediately reported to the IBC and OSP. Spills or accidents occurring in high containment (BL3 or BL4) laboratories resulting in an overt or potential exposure must be immediately reported to OSP.

Any spill or accident involving recombinant or synthetic DNA research of the nature described in the previous paragraph or that otherwise leads to personal injury or illness or to a breach of containment must be reported to the IBC and OSP. These kinds of events might include skin punctures with needles containing recombinant or synthetic DNA, the escape or improper disposition of a transgenic animal, or spills of high-risk recombinant materials occurring outside of a biosafety cabinet. Failure to adhere to the containment and biosafety practices articulated in the NIH Guidelines must also be reported to IBC and OSP.

In addition, exposure to Infectious Agents or Other Potentially Infected Material should be reported to the IBC.

Minor spills of low-risk agents not involving a breach of containment that were properly cleaned and decontaminated generally do not need to be reported. If the investigator, or other institutional staff are uncertain whether the nature or severity of the incident warrants reporting, contact the Biosafety Officer who can assist in making this determination, with guidance from OSP, if necessary.

Please complete the form below and submit to the Biosafety Officer (801-581-6590).

SPILL OR EXPOSURE EVENT REPORT

If a spill or exposure to Recombinant DNA or Synthetic Nucleic Acid Molecules, Infectious Agents or Other Potentially Infected Material occurs in your laboratory, please complete the following information and give to the Biosafety Officer. This will serve as a record of the event and be used for NIH OSP reporting, if necessary.

Date of Event:

Principal Investigator:

IBC Registration #:

Building/Room #:

Location with room (bench, biosafety cabinet, centrifuge, etc.):

Is this work NIH funded?

If yes, please provide the grant number, NIH funding institute or center, NIH program officer contact information (name, email etc):

Type of Event (Spill, puncture wound, etc.):

Names of lab personnel involved/exposed:

Site of exposure (eye, mouth, etc: If applicable):

Agent (Recombinant Lentivirus, recombinant adenovirus, etc);

Gene encoded by vector:

Known Oncogene or predicted to increase risk of oncogenesis (Y or N):

Containment Level (BSL-2-enhanced, BSL-2, etc)

Amount exposed (pfu, µl, etc: if known):

Narrative of post exposure/spill response (please describe disinfectant agents, washing of exposed sites, PPE worn during clean up, etc.)

Did the affected individuals need medical treatment and, if so, where were they treated?

Date submitted to Biosafety Officer:

Appendix 4 SOPs for Working with Human Cells

Minimum Personal Protective Equipment:

  - Face protection (goggles,  face shield or other splatter guard) is used to protect from potential aerosols, splashes or sprays of infectious or other hazardous materials to the face when cells must be taken outside the BSC (e.g. for microscopy).

  - Protective laboratory coats or gowns, designated for lab use are worn while in the laboratory. This protective clothing is removed and left in the laboratory before leaving for non-laboratory areas (e.g., cafeteria, library, administrative offices). All protective clothing is either disposed of in the laboratory or laundered by the institution; it should never be taken home by personnel.

  - Gloves are worn when hands may contact potentially infectious materials, contaminated surfaces or equipment. Gloves are disposed of when overtly contaminated, and removed when work with infectious materials is completed or when the integrity of the glove is compromised. Disposable gloves are not washed, reused, or used for touching "clean" surfaces (keyboards, telephones, etc.), and they should not be worn outside the lab. Alternatives to powdered latex gloves should be available. Hands are washed following removal of gloves.

Potential Hazards:

The potential laboratory hazards associated with human cells and tissues include the bloodborne pathogens HBV and HIV, as well as agents such as Mycobacterium tuberculosis that may be present in human lung tissues. Potential hazards to laboratory workers are presented by cells transformed with viral agents, such as SV-40, EBV, or HBV, as well as cells carrying viral genomic material. Tumorigenic human cells also are potential hazards as a result of self-inoculation.

Procedure:

Human cells should be handled using BIOSAFETY LEVEL 2 practices and containment. All work should be performed in a biosafety cabinet, and all material should be decontaminated by autoclaving or disinfection before discarding. All employees working with human cells and tissues should be enrolled in the institutional Bloodborne Pathogens Program, and should work under the policies and guidelines established by the institutions' Exposure Control Plan.

Employees who have not been vaccinated or have not received a recent HBV vaccination should be tested for immunological status against HBV and, if necessary, offered immunization, and be evaluated by a health care professional following an exposure incident.

Appendix 5 Model Lentivirus Protocol

- Preparation of Concentrated Lentiviral Particles

NOTE: All protocols using lentiviruses must be reviewed and approved by the Institutional Biosafety Committee. BSL-2 techniques are employed when generating and utilizing lentiviruses unless they encode oncogenes or toxins, or products that knock-down the expression of tumor suppressors: in such cases BSL-2-enhanced techniques must be employed

1.	Expand HEK 293T cells in DMEM media supplemented with 10% FBS (heat inactivated) in 10 cm tissue culture dishes. The cells are cultured at 37 °C in 5% CO2.

For unknown reasons, high passage HEK 293T cells do not produce high titer lentivirus. It is best to obtain a fresh stock of HEK 293Tfrom ATCC rather than acquire stocks from other laboratories. Also, cryopreserve early passage cells and thaw a fresh stock for the generation of high titer lentivirus.

- Aspirate the media and wash each plate with 5 ml of PBS.

- Add 1 ml/plate of trypsin and incubate for 2-5 min. Gently pipette up and down to mix the trypsin solution.

- Add 10 ml/plate of DMEM media supplemented with 10% FBS (heat inactivated) and determine the number of cells per ml using a hemocytometer.

- Plate 4-7 x 106 cells into each 10 cm tissue culture dish and adjust the volume of each plate to 10 ml with DMEM media supplemented with 10% FBS (heat inactivated). Culture the cells overnight at 37°C in 5% CO2.

Alternatively a confluent 10cm plate of HEK293T cells can be re-plated at a 1:10 dilution and cultured for 2-5 days until they are approximately 90% confluent. A typical procedure utilizing 24 10cm plates yields approximately 0.9-1.2 ml of a solution containing between 5.0 x107 to 5.0 x108 infectious units/ml (IU/ml).

- Aspirate the media and add 8 ml/plate of DMEM media supplemented with 10% FBS (heat inactivated) that has been pre-warmed to 37°C and culture for 2 h.

- Prepare PEI at 1 mg/ml (w/v) in deionized sterile water.

The transfection procedure described in this protocol is optimized for PEI (Huh et al., 2007). The PEI solutions can be aliquoted into 1.5 ml microcentrifuge tubes and stored at -20°C for up to a year. If a different transfection reagent will be used, optimization will be necessary.

- Add 720 µl of the 1 mg/ml (w/v) PEI solution to 11.28 ml of OPTI-MEM media in a 50 ml conical tube, which yields a solution sufficient to transfect 24 plates (30 µg of PEI per plate in 500 µl of OPTI-MEM).

The volume can be adjusted depending on the number of plates utilized.

- Prepare a solution of plasmids in a separate 15 ml conical tube, which is enough to transfect 24 plates, by adding 120 µg of pHIV-Luc-ZsGreen (5 µg/plate), 40.8 µg of pMDLg/pRRE (1.7 µg/plate), 40.8 µg of pRSV-Rev (1.7 µg/plate), and 40.8 µg of pCMV-VSV-G (1.7 µg/plate) to a final volume of 12 ml of OPTI-MEM media (500 µI/plate).

It is important to utilize plasmid solutions at >0.5 µg/µl since lower concentrations can dilute the media/DNA solution and decrease the transfection efficiency. In addition, to achieve high viral titers, it is necessary to mix all four plasmids together before mixing with the PEI solution.

- Add 12 ml of the plasmids/OPTI-MEM to the 12 ml of PEI/OPTI-MEM and mix. Incubate the solution for 30 min at room temperature.

- Add 1 ml/plate dropwise of the PEI and plasmid OPTI-MEM solution, gently swirl the plates and culture cells with solution for 24 h at 37 °C in 5% C02.

- Every time before working with solutions containing virus, place a large beaker containing >50 ml of bleach in a biosafety cabinet to disinfect any media and pipettes (make sure the final concentration of bleach is >10% before removing it from the biosafety cabinet).

Appropriate personal protection equipment are required. Tubes containing virus can only be opened in a biosafety cabinet. Sealed tubes must he sprayed with 70% ethanol solution before removal from the biosafety cabinet.

- Remove the media from each plate and transfer it to the beaker containing bleach. Add 8 ml/plate of DMEM media containing 10% FBS (heat inactivated) that has been pre-warmed to 37°C and incubate overnight at 37°C in 5% CO2.

- 48 h post transfection, remove the media from each plate and transfer it to a 50 ml conical tube on ice. Immediately add 8 ml/plate of fresh pre-warmed DMEM media containing 10% FBS (heat inactivated) to each plate and culture overnight at 37 °C in 5% CO2.

The transfection efficiency can be estimated at this stage by using a tissue culture microscope with fluorescence capability and observing the percent of cells fluorescent for Zsgreen. It is advisable that this is performed on an extra plate of cells that are fixed with 4% paraformaldehyde, to inactivate the lentivirus, prior to microscopy. Continue with the concentration procedure only if >50% of the cells are fluorescent. If <50% of the cells are transfected, optimize the transfection procedure before proceeding.

- Centrifuge the 50 ml conical tubes containing the media for 10 min at 2,500 x g at 4°C to pellet any large debris.

- Pass the supernatant through a 0.45 µm cellulose acetate filter and transfer 32 ml (about 4 plates worth) to each of the six ultracentrifuge tubes. Place tubes in titanium buckets on ice. Balance bucket pairs by adjusting the weight with DMEM media (this is done in the biosafety cabinet until the buckets are sealed).

- To pellet the virus, ultra-centrifuge the solution for 1 h 45 min at 112,000 x g at 4 °C using a swinging bucket rotor.

- Carefully remove the supernatant and transfer it to a beaker containing bleach. Immediately (do not let the pellet dry) add 150 µl of DMEM/F12 media to each centrifuge tube on ice. Gently pipette up and down without creating bubbles to dissolve the pellet. Place the ultracentrifuge tubes in a 50 ml conical tube and cap the conical tube to prevent evaporation of the solution. Store the conical tubes at 4 °C overnight.

- The following day, collect the second round of media from the plates (72 h post transfection) and process it in the same manner as steps 14-17.

Add bleach at a final concentration of 10% to the plates, incubate at room temperature for 20 minutes in the biosafety cabinet, and discard plates and media.

- After the ultracentrifugation step 17, remove the supernatant. Resuspend the pellets by transferring the DMEM/F12 solution containing the concentrated virus that was collected 48 h post transfection to each tube. Gently pipette up and down to dissolve the pellet. Place the ultracentrifuge tubes in a 50 ml conical vial to prevent evaporation and incubate at 4 °C for 2 h to ensure the virus is completely dissolved.

- Pipette up and down to resuspend the pellets, being careful to limit bubbles. Combine all of the resuspended pellets into one tube and gently pipet to mix. Aliquot 50-100 µl of the concentrated virus into 500 µl microfuge tubes and store at -80 °C.

The virus solutions are aliquoted to minimize freeze/thaws, which can significantly reduce the number of IU/ml. The virus solutions can be stored for up to 1 year at -80 °C.

- Determining Infectious Units (IU)lml by FACS

  - Expand HEK 293T cells in DMEM media supplemented with 10% FBS (heat inactivated) in 10 cm tissue culture dishes.

  - Aspirate the media and wash each plate with 5 ml of PBS.

  - Add 1 ml/plate of trypsin and incubate for 2-5 min at 37°C in 5% CO2. Gently pipette up and down to mix the trypsin solution.

  - Add 10 ml /plate of DMEM supplemented with 10% FBS (heat inactivated) and determine the number of cells per ml using a hemocytometer.

5.	Plate 2.0 x 105 cells in each well of a 6 well tissue culture plate and adjust the volume to 3 ml/well with DMEM media supplemented with 10% FBS (heat inactivated ). Culture the cells overnight at 37 °C in 5% CO2.

6.	Aspirate the media from four wells of the 6 well plate and ad d 1 ml/ well of DMEM media supplemented with 10% FBS ( heat inactivated) that has been pre-warmed to 37°C.

7.	Aspirate the media from the remaining two wells of the 6 well plate and carefully wash with 2 ml/well of PBS so the cells do not detach.

8.	Add 0.5 ml of trypsin to each of the two wells and incubate for 2-5 min at 37 °C in 5% CO2. Gently pipette up and down to mix the trypsin solution.

- Add 0.5 ml/well of DMEM media supplemented with 1 0% FBS and determine the number of cells present in each well independently using a hemocytometer.  Calculate the average number of cells present in each of the two wells. This number will be used to determine the IU/ml.

- Thaw an aliquot of concentrated virus in the biosafety cabinet.

- Dilute the virus solution 1:10 in DMEM medium supplemented with 10% FBS (heat inactivated). For example, add 8 µl of virus solution to 72 µl of media.

- Add different amounts of the diluted virus to each of the remaining four wells as described below, gently mix and culture overnight.

  - No virus control

  - 1 µl of diluted virus (1:10,000 final dilution)

  - 10 µI of diluted virus (1: 1,000 final dilution)

d) 	50 µI of diluted virus (1:200 final dilution)

The final dilution values are based upon the use of a 1:10 initial dilution of the virus and 1 ml of media in each well. Different dilutions are required since a sample containing <15% ZsGreen positive cells determined by FACS is required to more accurately calculate the IU/ml.

- Remove the media 24 h post infection and aspirate into bleach. Add 3 ml/well of DMEM media supplemented with 10% FBS that has been pre-warmed to 37°C and culture for 2 additional days.

- Remove the media 72 h post infection and aspirate into bleach. Carefully wash the wells with 2 ml/well of PBS so the cells do no detach.

- Add 0.5 ml/well of trypsin and incubate for 2-5 min. Gently pipette up and down to mix the trypsin solution.

- Transfer the cell suspension to a 15 ml conical tube on ice. Wash each well with 1 ml of HBSS containing 2% FBS and transfer this solution to the respective 15 ml conical vial.

- Add paraformaldehyde to a final concentration of 2% to each tube and incubate at 4 °C for 20 minutes to inactivate any residual virus that may be present.

- Centrifuge the cells for 5 min at 400 x g at 4 °C.

- Aspirate the supernatant and carefully resuspend the pellet with 1 ml of HBSS containing 2% FBS

- Centrifuge the cells for 5 min at 400 x g at 4 °C.

- Aspirate the supernatant and carefully resuspend the pellet with 250 µl of HBSS containing 2% FBS

- Pass the cell suspension through a 100 µm cell strainer and transfer to a FACS tube

- Setup the gates and detector voltages, and analyze samples using the sequential gates described below.

  - SSC vs. FSC: eliminate debris and gate only single cells

  - FLl (ZsGreen) histogram: determine percent ZsGreen positive cells

- Calculate the IU/ml using the formula below.

IU/ml = (average # of cells/well) * (% ZsGreen positive cells/100) *(virus dilution)

The average # of cells/well was determined above in step 9. We typically use a dilution where the cells are <15% Zs Green positive.  Titers ranging from 5 x 107 to 5 x 108 IU/ml are commonly achieved using this protocol.

- Transduction of cells

  - Seed 250,000 cells per well in a 6-well plate in 0.5 ml of medium.

For suspension infections, use ultra-low attachment plates, and for monolayer infections, use regular tissue culture-treated plates. The best method of infection must be determined empirically for each tumor. For monolayer infections, allow the cells to attach for 24-48 hr prior to adding virus.

  - Add polybrene to a final concentration of 1-4 µg/ml. The actual concentration should be empirically determined using the target cells. Add concentrated lentivirus using a multiplicity of infection (MOI) of 5-20 viral particles/target cell. Add additional medium to bring the final volume to 1 ml per well.

The multiplicity of infection may also have to be optimized for different cells.

  - Replace the virus-containing medium with fresh medium after overnight incubation.

  - Remove the virus medium and add 2 ml of fresh medium.

  - The cells should be handled using BSL-2 procedures until at least 72 hours has passed following infection and medium has been exchanged using at least five washes (50-Fold dilution with each wash).

  - Incubate at 37°C with 5% CO2.  Expression of virus proteins (e.g. ZsGreen) should be detectable by fluorescence microscopy after 72 hours. Feed the cells every day by exchanging the medium. Expand the cells until a sufficient number are obtained for transplantation.

- Transplantation of lentivirus-transduced into mice

  - Transduced cells should be considered potentially infectious and handled under BSL-2 procedures until they have undergone at least five washes (50-Fold dilution with each wash).

  - Transduced cells will not be transplanted until Step 1 is performed and at least 72 hours has passed following the infection procedure.

  - Cells should be transplanted under ABSL1 conditions following procedures approved by IACUC and by the IBC.

Appendix 6  Testing Lentiviral Vector Stocks for Replication Competent Virus

RCV Test Protocol:

- Harvest fresh vector-infected supernatant (vector stock)

- Clear cells from growth medium by low-speed centrifugation.

- Filter medium through 0.45 µm-pore-size cellulose acetate filters.

- Transduce SupT1* lymphocytes

  - Plate SupT1* cells into a 6 well plate at 105 cells/well in presence of 8 µg/ml of Polybrene (to enhance virus detection) and add 2 ml of RPMI-1640 Medium supplemented with 10% FBS (heat inactivated).

  - Infect SupT1* cells with serial dilutions of vector stock (starting with > 1 mL of unconcentrated supernatant or concentrated equivalent)

- Grow SupT1* cells for 3 weeks or longer (passage as necessary)

  - Harvest supernatant as per above in Step 1.

  - Perform ELISA assay for p24 antigen.

  - Record results for each batch (expressed as infectious units/mL).

# Several alternative sensitive end-point assays rather than the ELISA-based p24 Gag antigen assay have been described, including a product enhanced reverse transcriptase (PERT) assay that involves the vector’s reverse transcriptase, a PCR-based assay that detects Psi-Gag sequences from a recombination event between vector and packaging constructs, and PCR-based assays to detect the VSV-G Env used for pesudotyping. All represent acceptable methods for RCV testing.

Appendix 7  Viral Vector Maps

Appendix 8: Gene Editing Questionnaire

- Please specify which gene editing technology is being used (CRISPR/Cas9, Zinc Finger nucleases, TALENS, Meganucleases, other?).

- How is the gene editing technology being delivered to cells (e.g., nanoparticles, plasmid, lentivirus, adeno-associated virus, etc.)?

- What are the gene targets for these experiments?

- Are any of the targets oncogenes, tumor suppressors, or genes involved in cell cycle regulation?

- Are you targeting embryos or germ line cells?

- Are you planning research using gene editing technologies in humans, human embryos, or human germ cells?

(No gene editing of the germ line, human embryos or germ cells for clinical application is allowed. Gene editing of human embryos and germ cells for scientific purpose may be allowed, but must be evaluated on a case-by-case basis by the appropriate federal and local scientific review committees.)

- Please identify the cell types, tissues and/or organisms that are being targeted.

Answer the following questions for studies involving CRISPR/Cas9:

- Are the guide RNA (gRNA) and nuclease encoded on the same plasmid, viral vector or other delivery vehicle? If yes, describe.

- For gene editing research involving viral vectors, have potential off target effects by your gRNA been identified in the human genome?  This is helpful in assessing the risk of potential exposure in the event of an incident. The IBC recommends GT Scan (http://bioinformatics.csiro.au/gt-scan/), but other utilities may be used. Please identify which utility you have used and provide a list of off target genes (with up to 3 mismatches).

- Are you performing a “gene drive” experiment?

(A “gene drive” is established when Cas9 and gRNA are flanked by homology arms that match DNA sequences surrounding the genome of the gRNA target site. When such a system is introduced into the germ line of a sexually reproducing species, genomic alterations are continually transmitted to offspring.  (Reference: Akbari, Omar S., et al. "Safeguarding Gene Drive Experiments in the Laboratory." Science 349.6251 (2015): 927-9)?

- Does the construct encoding either Cas9 or gRNA also contain DNA with homology to genomic DNA that flanks the gRNA target site? If so, answer the questions below.

  - Will the experiment introduce this construct into the germ line of an organism?

  - Will the experiment result in a transgenic, sexually reproducing organisms?

  - If the answer is yes to both 11 and 11a, provide a complete description of your experiment.

Appendix 9: Biological Toxin SOP

[Customize text in parentheses and brackets to specific procedures and equipment in your laboratory.  Please refer to the Biosafety in Microbiological and Biomedical Laboratories, Appendix I: Guidelines for work with Toxin of Biological Origin for more information, https://www.cdc.gov/biosafety/publications/bmbl5/bmbl5_appendixi.pdf#x2013.]

| Standard Operating Procedures for [Toxin] |
| --- |
| #1 Chemicals/Hazards | [Obtain specific toxin hazard information from MSDS/SDS.] CAS number: [XXX] Routes of exposure: [XXX] How exposure might occur: [XXX] Target organs: [XXX] Signs/symptoms of exposure: [XXX] |
| #2 Prior to Work | Hazardous chemical and specific SOP training will be provided to personnel working with toxin and any other personnel authorized or required to be in the laboratory during toxin work. Appropriate inactivation method(s) for [toxin] will be determined and supplies for inactivation and spill cleanup of [toxin] will be readily available. [List vaccinations or antitoxins required or recommended for toxin]. If vaccinations or antitoxins are required, contact Occupational and Environmental Health and Safety at 801-581-6590. |
| #3 Environmental/ Ventilation Controls | Work with [toxin] will be performed in a [chemical fume hood/Biological Safety cabinet (BSC)]. List the type of BSC to be used (e.g. Class II, Type A2). In-line HEPA filters will be used on vacuum lines. Safety centrifuge cups or sealed rotors will be used if centrifuging materials containing [toxin], and the outside surfaces will be routinely decontaminated after each use. |
| #4 Personal Protective Equipment (PPE) | The following PPE will be worn when working with [toxin]: [Customize list] Laboratory coat or gown with long cuffed sleeves Disposable lab coat Disposable sleeves Safety glasses with side shields or chemical safety goggles Face protection such as a face shield if splash/spatter possible Gloves [type] that are impervious to [toxin] and diluent Respiratory protection [if aerosol hazard is present] If respirators are used the worker must be enrolled in the Respiratory protection program; contact OEHS for information (http://d2vxd53ymoe6ju.cloudfront.net/wp-content/uploads/sites/4/20160922155338/Resp-Prot-Program-rev.2016.pdf). Gloves must be changed immediately if contaminated, torn, or punctured. |
| #5 Special Handling Procedures & Storage Requirements | HANDLING Prep Sign will be posted on the room door when toxin is in use stating: “Toxins in Use -- Authorized Personnel Only.” All preparation of [toxin] will be performed over plastic-backed absorbent pads in a [fume hood/BSC]. Pads will be disposed of immediately upon contamination and after completion of tasks. Describe how toxin will be prepared: [Example: Vials of [toxin] will be purchased in pre-weighed powder form and then reconstituted in a [fume hood/biological safety cabinet (BSC)]. Weighing the [toxin] is not necessary as reconstitution will occur in the purchased vial and then aliquoted into vials with caps.] Use Only needle locking (Luer-Lock type) syringes or disposable syringe units will be used for injection or aspiration of [toxin]. A sharps container will be in the immediate vicinity for safe sharps disposal. Containers will be decontaminated before they are removed from [fume hood/BSC]. The [fume hood/BSC] will be decontaminated upon completion of tasks with [decontaminant and concentration] for [contact time]. All potentially contaminated disposable items will be placed in a hazardous waste bag and decontaminated before disposal. Hands will be washed upon completion of tasks. STORAGE [Toxin] will be stored in locked [freezer/refrigerator/cabinet/box/other] in [secure location room #]. TRANSPORT [Toxin] will be transported in labeled and sealed non-breakable secondary containers. |
| #6 Spill and Accident Procedures [Specific cleaning, decontamination agents (and contact times)/equipment and waste disposal procedures must be determined.] | All spills will be cleaned by properly protected and trained personnel only. Wash hands thoroughly after completing any spill clean-up. If you are not trained or comfortable cleaning up a spill, call OEHS for assistance at 801-581-6590. If it is an emergency (risk of exposure to others such as an on-going toxin release), call 911. Liquid spills: Personnel cleaning up a liquid spill will wear a lab coat/gown with cuffed sleeves (or disposable sleeves), goggles, and two pairs of nitrile gloves. Cover spill with absorbent paper towels and apply [inactivating agent + concentration], starting at the perimeter and working towards the center, allowing [XX min] contact time to deactivate [toxin]. Clean the spill area with [inactivating agent], then soap and water. The decontaminated spill waste will be double bagged and disposed of in the biohazard waste container. Powder spills inside of [fume hood/BSC]: Personnel cleaning up a powder spill will wear a lab coat/gown with cuffed sleeves (or disposable sleeves), goggles, and two pairs of nitrile gloves. Gently cover powder spill with dampened absorbent paper towels to avoid raising dust. Apply [inactivating agent + concentration], starting at the perimeter and working towards the center, allowing [XX min] contact time to deactivate [toxin]. Clean the spill area with [inactivating agent], then soap and water. The decontaminated spill waste will be double bagged and disposed of in the biohazard waste container. Powder spills outside of a [fume hood/BSC]: Remove all personnel from the room and restrict access; do not attempt to clean up the spill unless personnel are authorized to use a respirator. If personnel are not cleared to use a respirator, report the spill by notifying OEHS (at 801-581-6590). Tell them that a spill has occurred, and you need OEHS to assist with the spill cleanup. Be prepared to provide the following information: Name and phone number of knowledgeable person that can be contacted: [emergency contact name and phone number] [Toxin name], concentration and amount spilled, liquid or solid spill Number of injured, if any Location of spill This information can also be used in reporting to the Emergency Department after potential exposure. Personnel cleaning up a powder spill will wear a lab coat/gown with cuffed sleeves [or disposable sleeves], goggles, two pairs of nitrile gloves and a respirator. Gently cover powder spill with dampened absorbent paper towels to avoid raising dust. Apply [inactivating agent + concentration], starting at the perimeter and working towards the center, allowing [XX min] contact time to deactivate [toxin]. Clean the spill area with [inactivating agent], then soap and water. The decontaminated spill waste will be double bagged and disposed of in the biohazard waste container. For questions on spill cleanup, contact OEHS at 801-581-6590 for guidance. Wash hands thoroughly after completing any spill clean-up. |
| EXPOSURE PROCEDURES In Case of Emergency | Provide First Aid Immediately For sharps injury (needlestick or subcutaneous exposure), scrub exposed area thoroughly for 15 minutes using warm water and soap. For skin exposure, wash the area with soap and water. For large exposures, use the nearest safety shower for 15 minutes. Stay under the shower and remove clothing. Use a clean lab coat or spare clothing for cover-up. For eye exposure, use the eye wash for 15 minutes while holding eyelids open. For inhalation, move out of contaminated area. Get medical help. Get Help Call 911 or go to nearest Emergency Department (ED). Give details of exposure, i.e. agent, dose, route of exposure, time since exposure. Bring to the ED the MSDS/SDS and this SOP. Notify your supervisor as soon as possible for assistance. Secure area before leaving. Report Incident to Occupational and Environmental Health & Safety If serious accident, hospitalization or fatality, notify OEHS immediately after providing first aid and/or getting help. Call at 801-581-6590. |
| #7 Waste Disposal and Cleaning | Any waste [toxin] will be decontaminated or autoclaved as appropriate before disposal or given to OEHS for disposal whenever possible. Work space surfaces must be wiped down after completion of tasks with [inactivating agent + concentration] during the length of the experiment. Absorbent pads will be replaced after completion of tasks or immediately if contaminated. Used and potentially contaminated absorbent pads, PPE, etc. will be placed in a hazardous waste bag and autoclaved. If in-lab inactivation is not possible for [toxin] waste, it must be managed as hazardous chemical waste. Be aware that some form of treatment in the lab may be required before it can be managed as chemical waste. Contact OEHS at 801-581-6590 for disposal instructions. For chemical waste pick up complete a request through the lab management system (http://oehs.utah.edu/topics/lab-management-system). |
| #8 Special Precautions for Use of [Toxin] in Animals (This section must be completed if working with toxin in animals) Identify where animals will be injected, housed and any special precautions/warnings for animal handlers. | Use of toxins in animals will be documented and approved by IACUC. [Give detailed procedures for safely completing tasks, containment, decontamination information, and any special disposal requirements.] [Animals will be anesthetized or placed into a restraining apparatus before procedures using [toxin] are performed. Once the animal has been properly fitted into the restraining apparatus, the syringe will be loaded just prior to injection.] After procedures are complete, the restraining apparatus and surrounding work station will be decontaminated [inactivating agent + concentration]. All reusable lab equipment will be autoclaved. [Give any special disposal requirements] |
| #9 Approval Required | The protocol must be approved by the Institutional Biosafety Committee prior to commencement. All staff working with [toxin] must be trained on this SOP prior to starting work. They must also be trained on the [toxin] MSDS/SDS, and it must be readily available in the laboratory. All training must be documented and maintained by the PI. |
| #10 Decontamination | All surfaces will be decontaminated with [inactivating agent + concentration] after removing the plastic backed pads. All reusable lab equipment will be autoclaved. Note that some disinfecting agents may not deactivate [toxin]. |
| #11 Designated Area | All work with [toxin] must be done in a designated laboratory, work space and [fume hood/BSC]. Signage must be placed on door to room when [toxin] is used. This work will be conducted in [Room #] |

# Appendix 10:  Agent Summary Statement

The BMBL contains summary statements for many bacteria, viruses, and zoonotics.  If your agent is listed in the BMBL‘s Agent Summary Statements section (http://www.cdc.gov/biosafety/publications/bmbl5/index.htm), you may find it useful to paste it into the biosafety manual. Data on Salmonella typhimiruim is provided as an example

Salmonella typhimiruim

Salmonellae are gram-negative enteric bacteria associated with diarrheal illness in humans. They are motile oxidase-negative organisms that are easily cultivated on standard bacteriologic media, although enrichment and selective media may be required for isolation from clinical materials. Recent taxonomic studies have organized this genus into two species, S. enterica and S. bongori, containing more than 2500 antigenically distinct subtypes or serotypes. S. enterica contains the vast majority of serotypes associated with human disease. S. enterica serotypes Typhimurium and Enteritidis (commonly designated S. Typhimurium and S. Enteritidis) are the serotypes most frequently encountered in the United States. This summary statement covers all pathogenic serotypes except S. Typhi.

Occupational Infections

Salmonellosis is a documented hazard to laboratory personnel. Primary reservoir hosts include a broad spectrum of domestic and wild animals, including birds, mammals, and reptiles, all of which may serve as a source of infection to laboratory personnel. Case reports of laboratory-acquired infections indicate a presentation of symptoms (fever, severe diarrhea, abdominal cramping) similar to those of naturally-acquired infections, although one case also developed erythema nodosum and reactive arthritis.

Natural Modes of Infection

Salmonellosis is a food borne disease of worldwide distribution. An estimated 5 million cases of salmonellosis occur annually in the United States. A wide range of domestic and feral animals (poultry, swine, rodents, cattle, iguanas, turtles, chicks, dogs, cats) may serve as reservoirs for this disease, as well as humans. The most common mode of transmission is by ingestion of food from contaminated animals or contaminated during processing. The disease usually presents as an acute enterocolitis, with an incubation period ranging from 6 to 72 hours.

Laboratory Safety and Containment Recommendations

The agent may be present in feces, blood, urine, and in food, feed, and environmental materials. Ingestion or parenteral inoculation are the primary laboratory hazards. The importance of aerosol exposure is not known. Naturally or experimentally infected animals are a potential source of infection for laboratory and animal care personnel, and for other animals

Strict compliance with BSL-2 practices, containment equipment, and facilities are recommended for all activities utilizing known or potentially infectious clinical materials or cultures. This includes conducting procedures with aerosol or high splash potential in primary containment devices such as a BSCs or safety centrifuge cups. Personal protective equipment should be used in accordance with a risk assessment, including splash shields, face protection, gowns, and gloves. The importance of proper gloving techniques and frequent and thorough hand washing is emphasized. Care in manipulating faucet handles to prevent contamination of cleaned hands or the use of sinks equipped with remote water control devices, such as foot pedals, is highly recommended. Special attention to the timely and appropriate decontamination of work surfaces, including potentially contaminated equipment and laboratory fixtures, is strongly advised. ABSL-2 facilities and practices are recommended for activities with experimentally infected animals.

Special Issues

Transfer of Agent Importation of this agent may require CDC and/or USDA importation permits. Domestic transport of this agent may require a permit from USDA/APHIS/VS. A DoC permit may be required for the export of this agent to another country. See Appendix C of the BMBL for additional information.

# Appendix 11: Post-Exposure Fact Sheet

This is an example of a post exposure procedure for working with pathogenic organisms: edit as necessary. Relevant information can be found in Pathogen Safety Data Sheets (PSDSs), technical documents that describe the hazardous properties of a human pathogen and recommendations for work involving these agents in a laboratory setting. These documents have been produced by the Public Health Agency of Canada (the Agency) as educational and informational resources for laboratory personnel working with these infectious substances. PSDSs can be found at http://www.phac-aspc.gc.ca/lab-bio/res/psds-ftss/index-eng.php

Post-Exposure Procedures:

Salmonella typhimiruim

Characteristics:  Family Enterobacteriaceae; gram negative rod; motile, aerobic and facultatively anaerobic; serological identification of somatic and flagellar antigens; over 2000 serotypes capable of causing disease.

Incubation Period:  Six to 72 hours, usually about 12-36 hours

Symptoms:   Salmonellosis is an acute gastroenteritis; acute infectious disease with sudden onset of abdominal pain, diarrhea, nausea and vomiting; may progress to more serious septicemia, includes focal infections, abscesses, endocarditis, pneumonia; may also cause typhoid like enteric fever; some cases develop reactive arthritis (Reiter's syndrome) which may become chronic

Infectious Dose: 100 - 1,000 organisms - ingestion; varies with multiple factors

What is a potential exposure? Ingestion, needlestick or cut with contaminated sharp object, splash to the eye, contact with broken skin.

Post-Exposure Treatment:  Skin exposure / Percutaneous:  Wash affected area with antibacterial soap and water (apply antiseptic (3% H2O2) if available), report to the Occupational Medical Clinic at the Redwood Health Center. Mucous membrane exposure (splash to eye): flush eyes for 15 minutes using eyewash, then report to Occupational Medical Clinic.  Ingestion: Report to the Occupational Medical Clinic.  Antibiotic therapy may be required.   If the Health Center is closed, go to the University of Utah Hospital Emergency Room.

If symptoms appear with no known incidence of exposure:  Seek medical attention and inform the health care provider of the microorganisms used in the workplace.

Prevention:  Biosafety level 2 practices, containment equipment, and facilities; wear lab coat and gloves; frequent handwashing is essential.  Never eat in the lab.  Caution should be used with sharps.

Reporting: Make note of the date and time of the incident and any relevant details.  Inform principal investigator, fill out an Incident Report Form (https://ehs.utah.edu/incident-report), and contact OEHS (1-6590).  If recombinant, the incident must be reported to NIH Office of Biotechnology Activities.

Appendix 12:  Use and Disposal of Sharps

To prevent needle stick injuries:

- Do not recap needles by hand.  RECAPPING OF NEEDLES IS PROBHIBETED.

- Do not remove needles from syringes by hand.

- Do not bend, break, or otherwise manipulate needles by hand.

- Avoid using needles whenever possible.

- Replace glass materials with plastic (such as Pasteur pipettes)

- Immediately after use, discard needle and syringe (whether contaminated or not) into puncture resistant sharps containers. RECAPPING OF NEEDLES IS PROHIBITED.

- Use a Food and Drug Administration (FDA)-cleared sharps container if you generate sharps waste (pictured below). A description of FDA-Cleared Sharps containers can be found here.  FDA-cleared sharps disposal containers are made from rigid plastic, come marked with a line that indicates when the container should be considered full, which means it’s time to dispose of the container, and have the Universal Biohazard symbol.

- Never discard sharps into regular trash.

- Never discard sharps into bags of biological waste.

- Use care and caution when cleaning up after procedures that require the use of syringes and needles.

- Do not overfill sharps containers.  Close completely when 3/4 full, request pickup from the OEHS through the Lab Management System (LMS) webpage at https://ehs.utah.edu/research-safety/lab-management-system.

- Locate sharps containers in areas in which needles are commonly used.  Make containers easily accessible.

- Replacement sharps containers may be obtained through the LMS or can be from laboratory supply distributors such as VWR and ThermoFisher.  Be sure to select sharps containers that withstand autoclaving.

In the event of a needle stick injury:

Wash the area thoroughly with soap and water.  Notify supervisor immediately, report to the Occupational Medical Clinic at the Redwood Health Center and fill out an Incident Report Form, as soon as possible: see Appendix 2.

Contaminated Serological Pipets and Pipet Tips

Serological pipets (glass and plastic) and disposable pipet tips are considered puncture hazards and should be disposed of as sharps.  BSL2 contaminated pipets and tips should be discarded in approved sharps containers, as described above.

Due to the large size of serological pipets, investigators disposing of large numbers of these can request 20 gallon hard-sided biohazard waste containers (pictured below) from OEHS through the LMS.  These will be picked up by OEHS staff as for other biohazardous waste.

20 Gallon Waste Container

Decontaminated Serological Pipets and Pipet Tips

It is possible to decontaminate serological pipets and tips prior to disposal.  Ensure that both the inside and outside of the pipets or tips are exposed to the approved disinfectant (e.g. a freshly prepared 1:10 dilution of bleach) for at least 20 minutes. However, serological pipets and disposable tips are still considered puncture hazards.  Therefore, after removing the disinfectant, they can be disposed of in a Broken Glass box (rigid puncture resistant boxes lined with a plastic bag and labeled “Broken Glass”: pictured below), which can be obtained from your custodial staff or from OEHS.  Once they are 2/3 full they should closed with tape and disposed as regular trash by your custodians.

Broken Glass Box

Appendix 13: Sharps injury Log

# Sharps Injury Log

##### Laboratory Name:________________		Year 2_____

| Date | Case/Report # | Type of Device (e.g. syringe, suture needle) | Brand Name of Device | Work Area where injury occurred (e.g. Geriatrics, Lab) | Brief description of how incident occurred (i.e., procedure being done, action being performed (disposal, injection, etc.), body part injured) |
| --- | --- | --- | --- | --- | --- |

29 CRF 1910.1030, OSHA’s Blood borne Pathogens Standard, in paragraph (h)(5), requires an employer to establish and maintain a Sharps Injury Log for recording all percutaneous injuries in a facility occurring from contaminated sharps. The purpose of the Log is to aid in the evaluations of devices being used in healthcare and other facilities and to identify problem devices or procedures requiring additional attention or review. This log must be kept in addition to the injury and illness log required by 29 CFR 1904. The Sharps Injury Log should include all sharps injuries occurring in a calendar year. The log must be retained for five years following the end of the year to which it relates. The Log must be kept in a manner that preserves the confidentiality of the affected employee.

Appendix 14: Disposal of Biohazardous Waste

Biowaste Disposal – Solids

The Department of Occupational and Environmental Health and Safety (OEHS) Lab Management System (LMS) allows research investigators to request hazardous material pickups by OEHS staff and request empty containers.  Please visit the LMS website for information.

Waste containers obtained from OEHS are solid sided, leak proof, lined with red biohazard bags, and labeled with a biohazard symbol.  Keep the container lid closed unless someone is working nearby and regularly adding waste to the container.

When the red bag is ¾ full, loosely tie or tape the bag closed. Secure the lid on the waste container and move it to a convenient storage location or transport it to a biohazardous waste storage room, if available.  Biohazardous waste must be moved or transported inside a rigid, leak- resistant, labeled container with the lid closed. Request a pickup from your lab using the LMS.

If you have an autoclave available for disinfection of biohazardous waste, place a red biohazard bag in a solid puncture resistant container.  Place a Ziploc bag or balloon containing water in the bag when it is about half full to generate steam during autoclaving. When the red bag is full, tie or tape the bag closed. Secure the lid on the waste container and move it to the autoclave room.

The bag should be removed and placed in a solid autoclave resistant tray: the bag should NEVER be placed directly on the floor. After the cycle, the bag may be disposed of as regular trash: indicators that the contents have been autoclaved must be present.

Biowaste Disposal – Liquids

Blood, aspirated tissue culture media, or other liquid waste generated from BSL2 experiments must be disinfected and then disposed. Bleach is typically used to disinfect liquids, but other agents, such as Wescodyne, may be used if effective.

If you use bleach:

- Ensure the final concentration exceeds 0.5% sodium hypochlorite (no less than one part bleach to 9 parts liquid).

- Ensure the bleach is fresh: in tissue culture media traps change at least twice weekly.

- Ensure the media is exposed to disinfectant for at least 20 minutes prior to disposal.

- Dispose down the sink

If you use Wescodyne:

- Ensure the final concentration exceeds 1% (no less than one part Wescodyne to 99 parts liquid).

- In tissue culture media traps change at least every 3 months (indicate the date of the last change on the flask).

- Ensure the media is exposed to disinfectant for at least 20 minutes prior to disposal.

- Collect waste into containers marked “Unwanted Materials” and date when you start collecting.  When full or 6 months after your start date (whichever happens first), arrange pickup by OEHS through the LMS website.  NO DRAIN DISPOSAL.

If the container will be unattended (outside of your immediate control) then label it with the date, time and the words “Biohazardous liquid” and keep it in a secondary container (for example, a plastic tub) while it is disinfecting.

If you use other agents to decontaminate liquid cultures follow the instructions on the packaging.  Contact the Biosafety Officer (801-581-6590) for advice on appropriate disinfectants and procedures for disposal of treated waste.

Mixed liquid and solid biohazardous waste

Mixed liquid and solid waste should be separated in a biosafety cabinet (decant the liquid from the solid). Manage the liquids and solids separately as detailed above.

State of Utah regulations for using Autoclaves for Treating Infectious Waste

Below are sections of the State of Utah Regulations for treating infectious waste that apply to the use of autoclaves.  These regulations “R315.  Environmental Quality, Waste Management and Radiation Control, Waste Management” came into effect April 2013.

R315-316-5.  Infectious Waste Treatment and Disposal Requirements.

(4)  Infectious waste may be sterilized by heating in a steam sterilizer to render the waste non-infectious.

(a)  The operator shall have available, and shall certify in writing that he understands, written operating procedures for each steam sterilizer, including time, temperature, pressure, type of waste, type of container, closure of container, pattern of loading, water content, and maximum load quantity.

(b)  Infectious waste shall be subjected to sufficient temperature, pressure and time to inactivate Bacillus stearothermophilus spores in the center of the waste load at a 6 Log10 reduction or greater.

(c)  Unless a steam sterilizer is equipped to continuously monitor and record temperature and pressure during the entire length of each sterilization cycle, each package of infectious waste to be sterilized shall have a temperature-sensitive tape or equivalent test material, such as chemical indicators, attached that will indicate if the sterilization temperature and pressure have been reached.  Waste shall not be considered sterilized if the tape or equivalent indicator fails to indicate that a temperature of at least 250 degrees Fahrenheit (121 degrees Celsius) was reached and a pressure of at least 15 psi was maintained during the process.

(d)  Each sterilization unit shall be evaluated for effectiveness with spores of B. stearothermophilus at least once each 40 hours of operation or each week, whichever is less frequent.

(e)  A written log for each load shall be maintained for each sterilization unit which shall contain at a minimum:

(i)  the time of day and the date of each load and the operator's name;

(ii)  the amount and type of infectious waste placed in the sterilizer; and

(iii)  the temperature, pressure, and duration of treatment.

Appendix 15: Safe Handling of Cryogenic Liquids

Danger!  Vials immersed in liquid nitrogen may explode violently when removed!  Wear face and eye protection!   Plastic vials (even Nunc vials with silicon O-rings) used for storing cells in liquid nitrogen are designed to be used in the liquid nitrogen vapor phase.  When immersed in the liquid phase, the liquid nitrogen frequently enters vials around the cold O-ring.  When vials are removed to room temperature, the liquid nitrogen in the vial immediately begins to boil.  Usually it escapes harmlessly past the seal.  Occasionally (about 1 out of 1000 vials), the seal is too tight, and the pressure causes a violent rupturing of the vial, sending shards of sharp plastic rocketing in unpredictable directions with sufficient energy to lacerate the face and cause severe eye injury.  When removing vials from liquid nitrogen, it is mandatory that you wear full face shields, pulled in to touch your chin so that shards can't fly under the shield.  If they fit, wear goggles underneath the face shield.

Appendix 16: BMBL BSL2 Laboratory Criteria and Design

Biosafety in Microbiological and Biomedical Laboratories (BMBL)

5th Edition, February 2007

Centers for Disease Control and Prevention and National Institutes of Health

Biosafety Level 2 builds upon BSL-1. BSL-2 is suitable for work involving agents that pose moderate hazards to personnel and the environment. It differs from BSL-1 in that:

- laboratory personnel have specific training in handling pathogenic agents and are supervised by scientists competent in handling infectious agents and associated procedures;

- access to the laboratory is restricted when work is being conducted; and

- all procedures in which infectious aerosols or splashes may be created are conducted in BSCs or other physical containment equipment.

The following standard and special practices, safety equipment, and facility requirements apply to BSL-2:

Standard Microbiological Practices

1. The laboratory supervisor must enforce the institutional policies that control access to the laboratory.

2. Persons must wash their hands after working with potentially hazardous materials and before leaving the laboratory.

3. Eating, drinking, smoking, handling contact lenses, applying cosmetics, and storing food for human consumption must not be permitted in laboratory areas. Food must be stored outside the laboratory area in cabinets or refrigerators designated and used for this purpose.

4. Mouth pipetting is prohibited; mechanical pipetting devices must be used.

5. Policies for the safe handling of sharps, such as needles, scalpels, pipettes, and broken glassware must be developed and implemented. Whenever practical, laboratory supervisors should adopt improved engineering and work practice controls that reduce risk of sharps injuries.

Precautions, including those listed below, must always be taken with sharp items. These include:

- Careful management of needles and other sharps are of primary importance. Needles must not be bent, sheared, broken, recapped, removed from disposable syringes, or otherwise manipulated by hand before disposal.

- Used disposable needles and syringes must be carefully placed in conveniently located puncture-resistant containers used for sharps disposal.

- Non-disposable sharps must be placed in a hard walled container for transport to a processing area for decontamination, preferably by autoclaving.

- Broken glassware must not be handled directly. Instead, it must be removed using a brush and dustpan, tongs, or forceps. Plastic ware should be substituted for glassware whenever possible.

6. Perform all procedures to minimize the creation of splashes and/or aerosols.

7. Decontaminate work surfaces after completion of work and after any spill or splash of potentially infectious material with appropriate disinfectant.

8. Decontaminate all cultures, stocks, and other potentially infectious materials before disposal using an effective method. Depending on where the decontamination will be performed, the following methods should be used prior to transport:

- Materials to be decontaminated outside of the immediate laboratory must be placed in a durable, leak proof container and secured for transport.

- Materials to be removed from the facility for decontamination must be packed in      accordance with applicable local, state, and federal regulations.

9. A sign incorporating the universal biohazard symbol must be posted at the entrance to the laboratory when infectious agents are present. Posted information must include: the laboratory’s biosafety level, the supervisor’s name (or other responsible personnel), telephone number, and required procedures for entering and exiting the laboratory. Agent information should be posted in accordance with the institutional policy.

10. An effective integrated pest management program is required. See Appendix G.

11. The laboratory supervisor must ensure that laboratory personnel receive appropriate training regarding their duties, the necessary precautions to prevent exposures, and exposure evaluation procedures. Personnel must receive annual updates or additional training when procedural or policy changes occur. Personal health status may impact an individual’s susceptibility to infection, ability to receive immunizations or prophylactic interventions. Therefore, all laboratory personnel and particularly women of child-bearing age should be provided with information regarding immune competence and conditions that may predispose them to infection. Individuals having these conditions should be encouraged to self-identify to the institution’s healthcare provider for appropriate counseling and guidance.

Special Practices

1. All persons entering the laboratory must be advised of the potential hazards and meet specific entry/exit requirements.

2. Laboratory personnel must be provided medical surveillance and offered appropriate immunizations for agents handled or potentially present in the laboratory.

3. Each institution must establish policies and procedures describing the collection and storage of serum samples from at-risk personnel.

4. A laboratory-specific biosafety manual must be prepared and adopted as policy. The biosafety manual must be available and accessible.

5. The laboratory supervisor must ensure that laboratory personnel demonstrate proficiency in standard and special microbiological practices before working with BSL-2 agents.

6. Potentially infectious materials must be placed in a durable, leak proof container during collection, handling, processing, storage, or transport within a facility.

7. Laboratory equipment should be routinely decontaminated, as well as, after spills, splashes, or other potential contamination.

- Spills involving infectious materials must be contained, decontaminated, and cleaned up by staff properly trained and equipped to work with infectious material.

- Equipment must be decontaminated before repair, maintenance, or removal from the laboratory.

8. Incidents that may result in exposure to infectious materials must be immediately evaluated and treated according to procedures described in the laboratory biosafety safety manual. All such incidents must be reported to the laboratory supervisor. Medical evaluation, surveillance, and treatment should be provided and appropriate records maintained.

9. Animals and plants not associated with the work being performed must not be permitted in the laboratory.

10. All procedures involving the manipulation of infectious materials that may generate an aerosol should be conducted within a BSC or other physical containment devices.

Safety Equipment (Primary Barriers and Personal Protective Equipment)

1. Properly maintained BSCs (preferably Class II), other appropriate personal protective equipment, or other physical containment devices must be used whenever:

- Procedures with a potential for creating infectious aerosols or splashes are conducted. These may include pipetting, centrifuging, grinding, blending, shaking, mixing, sonicating, opening containers of infectious materials, inoculating animals intranasally, and harvesting infected tissues from animals or eggs.

- High concentrations or large volumes of infectious agents are used.  Such materials may be centrifuged in the open laboratory using sealed rotor heads or centrifuge safety cups.

2. Protective laboratory coats, gowns, smocks, or uniforms designated for laboratory use must be worn while working with hazardous materials. Remove protective clothing before leaving for non-laboratory areas (e.g., cafeteria, library, administrative offices). Dispose of protective clothing appropriately, or deposit it for laundering by the institution. It is recommended that laboratory clothing not be taken home.

3. Eye and face protection (goggles, mask, face shield or other splatter guard) is used for anticipated splashes or sprays of infectious or other hazardous materials when the microorganisms must be handled outside the BSC or containment device. Eye and face protection must be disposed of with other contaminated laboratory waste or decontaminated before reuse. Persons who wear contact lenses in laboratories should also wear eye protection.

4. Gloves must be worn to protect hands from exposure to hazardous materials. Glove selection should be based on an appropriate risk assessment. Alternatives to latex gloves should be available. Gloves must not be worn outside the laboratory. In addition, BSL-2 laboratory workers should:

- Change gloves when contaminated, integrity has been compromised, or when otherwise necessary. Wear two pairs of gloves when appropriate.

- Remove gloves and wash hands when work with hazardous materials has been completed and before leaving the laboratory.

- Do not wash or reuse disposable gloves. Dispose of used gloves with other contaminated laboratory waste. Hand washing protocols must be rigorously followed.

5. Eye, face and respiratory protection should be used in rooms containing infected animals as determined by the risk assessment.

Laboratory Facilities (Secondary Barriers)

1. Laboratory doors should be self-closing and have locks in accordance with the institutional policies.

2. Laboratories must have a sink for hand washing. The sink may be manually, hands-free, or automatically operated. It should be located near the exit door.

3. The laboratory should be designed so that it can be easily cleaned and decontaminated. Carpets and rugs in laboratories are not permitted.

4. Laboratory furniture must be capable of supporting anticipated loads and uses. Spaces between benches, cabinets, and equipment should be accessible for cleaning.

- Bench tops must be impervious to water and resistant to heat, organic solvents, acids, alkalis, and other chemicals.

- Chairs used in laboratory work must be covered with a non-porous material that can be easily cleaned and decontaminated with appropriate disinfectant.

5. Laboratory windows that open to the exterior are not recommended. However, if a laboratory does have windows that open to the exterior, they must be fitted with screens.

6. BSCs must be installed so that fluctuations of the room air supply and exhaust do not interfere with proper operations. BSCs should be located away from doors, windows that can be opened, heavily traveled laboratory areas, and other possible airflow disruptions.

7. Vacuum lines should be protected with High Efficiency Particulate Air (HEPA) filters, or their equivalent. Filters must be replaced as needed. Liquid disinfectant traps may be required.

8. An eyewash station must be readily available.

9. There are no specific requirements on ventilation systems. However, planning of new facilities should consider mechanical ventilation systems that provide an inward flow of air without recirculation to spaces outside of the laboratory.

10. HEPA filtered exhaust air from a Class II BSC can be safely re-circulated back into the laboratory environment if the cabinet is tested and certified at least annually and operated according to manufacturer’s recommendations. BSCs can also be connected to the laboratory exhaust system by either a thimble (canopy) connection or a direct (hard) connection. Provisions to assure proper safety cabinet performance and air system operation must be verified.

11. A method for decontaminating all laboratory wastes should be available in the facility (e.g., autoclave, chemical disinfection, incineration, or other validated decontamination method).

# Footnotes

IBC requires review at least every 12 months

PI and IBC should conduct a risk assessment and determine if respiratory protection is required

IBC requires review at least every 12 months